## Slide 1: X-Ray Image-Based Navigation for Hip Osteotomy
Checkpoint Presentation
Mentors: |  | Dr. Mehran Armand, Dr. Yoshito Otake, Ryan Murphy
Course: |  | Computer Integrated Surgery II (EN.600.646)
Jesse Hamilton and Michael Van Maele
Group #3
March 29, 2012

## Slide 2: Summary
Summary | Dependencies | Milestones | Deliverables | Experiment | Results
Project Overview

## Slide 3: Summary
Current system: Periacetabular osteotomy (PAO) to correct developmental dysplasia of the hip (DDH) navigated by an optical tracker.
Summary | Dependencies | Milestones | Deliverables | Experiment | Results
Courtesy of Ryan Murphy
Optical tracker PAO setup.
Polaris camera.
Dynamic rigid bodies (x3)
Point-based registration.

## Slide 4: Summary
Current system: The Biomechanical Guidance System (BGS) utilizes angle and contact pressure to estimate the optimal fragment realignment .
Summary | Dependencies | Milestones | Deliverables | Experiment | Results
Courtesy of Ryan Murphy
3D model of hip and fragment.
PAO procedure.*
*http://www.hipandpelvis.com/patient_education/periace/page2.html

## Slide 5: Summary
Our goal: We want to replace optical tracker navigation with X-ray navigation.
Summary | Dependencies | Milestones | Deliverables | Experiment | Results
C-arm fluoroscopy.*
*http://www.anteriorhipreview.com/wp-content/uploads/2011/07/c-arm-during-anterior-total-hip-arthroplasty.jpg
Courtesy of Ryan Murphy
Optical tracker PAO setup.

## Slide 6: Summary
Our motivation:
  X-ray navigation is more in line with what surgeons are already doing.
  C-arm more accessible than optical tracker.
  Surgeons have occasionally disputed the results of the optical tracker method.
Summary | Dependencies | Milestones | Deliverables | Experiment | Results

## Slide 7: Summary
Our aims:
  to design and implement a surgical pipeline for X-ray image-based navigation for hip osteotomy, and
  to compare the proposed method with the current method (optical tracker-based navigation) experimentally.
Summary | Dependencies | Milestones | Deliverables | Experiment | Results

## Slide 8: Dependencies
Summary | Dependencies | Milestones | Deliverables | Experiment | Results
Descriptions & Status

## Slide 9: Dependencies
Summary | Dependencies | Milestones | Deliverables | Experiment | Results
Dependency | Target Date | Comment
Get mock OR access. | DONE | J-Cards enabled.
Schedule mentor meetings. | DONE | Thursdays at 10:30 AM, weekly.
Get radiation badges. | DONE | Stored in mock OR.
Get radiation training. | n/a | No response from training liaison.*
Get BGS-compatible PCs. | DONE | Hackerman B08
Install BGS software. | DONE | Hackerman B08
*Able to work with C-arm under mentor supervision.
No critical dependencies remain.

## Slide 10: Milestones
Summary | Dependencies | Milestones | Deliverables | Experiment | Results
Progress Report

## Slide 11
ORIGINAL SCHEDULE
Summary | Dependencies | Milestones | Deliverables | Experiment | Results

## Slide 12: Improved Schedule
Summary | Dependencies | Milestones | Deliverables | Experiment | Results

## Slide 13: Milestone 0
Convert high-level pipeline to low-level MATLAB script and debug with test data. (3/12)
Status: COMPLETED
Tasks:
  Perform a walk-through of the optical tracker-navigated procedure.
  Agree on pipeline flowchart for X-ray-navigated procedure.
  Convert chart components to script functions.
  Debug script functions with test data.
Summary | Dependencies | Milestones | Deliverables | Experiment | Results

## Slide 14: Milestone 1
Perform and report on mock surgery using X-ray navigation pipeline. (3/30)
Status: NEARLY COMPLETED
Summary | Dependencies | Milestones | Deliverables | Experiment | Results
Task | Completion | Comment | Target Date
Perform mock surgery. | 100% | In mock OR with pelvic phantom. | DONE
BB classification. | 100% | Classification is currently manual. | DONE
Code and debug pipeline in MATLAB. | 80% | Mild debugging issues. | 4/3
Write report on mock surgery. | 70% | Need fully formed pipeline to process data. | 4/3

## Slide 15: Milestone 1
Summary | Dependencies | Milestones | Deliverables | Experiment | Results
Report must be finished.

## Slide 16: Milestones 2 & 3
Prepare the navigation pipeline for online use. (4/13)
Status: BEGIN NEXT WEEK
Perform and report on cadaver surgery using X-ray navigation pipeline and conclude any pipeline adjustments. (5/4)
Status: BEGIN NEXT WEEK
Summary | Dependencies | Milestones | Deliverables | Experiment | Results

## Slide 17: Deliverables
Summary | Dependencies | Milestones | Deliverables | Experiment | Results
Full List and Progress Report

## Slide 18: Deliverables
Summary | Dependencies | Milestones | Deliverables | Experiment | Results
On schedule to achieve all deliverables (as of 3/29).
Minimum
  Manual BB classification solution. (3/16) (100%)
  X-ray navigation pipeline/protocol. (3/30) (80%)
  Report on results of mock surgery with pelvic phantom. (3/30) (70%)
Expected
  BGS software with X-ray navigation pipeline integrated. (4/20)
  Analysis of non-rigid FTRAC attachment methods. (4/20)
  Report on results of cadaver surgery. (4/27)
Maximum
  Simulated error analysis of FTRAC. (5/4)
  Automatically initialized EM pose estimation algorithm. (5/4)

## Slide 19: Nearly Completed Deliverable: Pipeline
Summary | Dependencies | Milestones | Deliverables | Experiment | Results

## Slide 20: Achieved Deliverable: BBs
BB Attachment:
  Trivial to do for pelvic phantom.
  May be trickier for the cadaveric specimen.
Summary | Dependencies | Milestones | Deliverables | Experiment | Results
Tear from BB reel.
Place on bone.
View with C-arm.

## Slide 21
Summary | Dependencies | Milestones | Deliverables | Experiment | Results
Experiment
Test of Pipeline with Pelvic Phantom
Notes: tested x-ray nav & optical methods, mock OR, plastic phantom

## Slide 22
Summary | Dependencies | Milestones | Deliverables | Experiment | Results
Experiment
Purpose:
  Demonstrate proof of concept for proposed x-ray navigation pipeline using:
    non-rigid FTRAC attachment
    pelvic phantom
  Acquire raw C-arm images to use when creating MATLAB code for pipeline.
  Gauge accuracy of X-ray navigation method versus optical tracker method.
Notes: Explain non-rigid vs. rigid FTRAC attachment
Optical method = gold standard

## Slide 23
Summary | Dependencies | Milestones | Deliverables | Experiment | Results
Attach “confidence” BBs to ilium
Attach “fragment” BBs
Place FTRAC within C-arm field of view
Acquire two “BEFORE” images
Acquire two “AFTER” images
Remove FTRAC. Realign bone fragment
Place FTRAC within C-arm field of view
Intraoperative Pipeline

## Slide 24
Summary | Dependencies | Milestones | Deliverables | Experiment | Results
Attach “confidence” BBs to ilium
Attach “fragment” BBs
Place FTRAC within C-arm field of view
Acquire two “BEFORE” images
Acquire two “AFTER” images
Remove FTRAC. Realign bone fragment
Place FTRAC within C-arm field of view
Confidence BBs
Fragment BBs
Intraoperative Pipeline
FTRAC
Notes: confidence BBs  virtual reference frame (fixed) on pelvis; lets us calculate fragment movement even when pelvis moves in between pre-op and post-op C-arm shots
highlight non-rigid “attachment” of FTRAC (difficult to line up in field of view  will look for better solution)
One of our fragment BBs fell off! Okay because 3 BBs are enough to do a registration (although less accurately)

## Slide 25
Summary | Dependencies | Milestones | Deliverables | Experiment | Results
Attach “confidence” BBs to ilium
Attach “fragment” BBs
Place FTRAC within C-arm field of view
Acquire two “BEFORE” images
Acquire two “AFTER” images
Remove FTRAC. Realign bone fragment
Place FTRAC within C-arm field of view
Intraoperative Pipeline
Notes: Acquring pre-op images
Rigid body is leftover from optical tracking

## Slide 26
Summary | Dependencies | Milestones | Deliverables | Experiment | Results
Imaging Subroutine
Segment BBs globally
Classify BBs as
FTRAC
Confidence
Fragment
Estimate C-arm pose using FTRAC
Register C-arm images with CT volume
Display biomechanics & radiographic angles
Dewarp C-arm distortion
Calibration phantom
Notes: Ofri phantom = grid with lines we place on C-arm
fits Bernstein polynomials (CIS I!)

## Slide 27
Summary | Dependencies | Milestones | Deliverables | Experiment | Results
Imaging Subroutine
Segment BBs globally
Classify BBs as
FTRAC
Confidence
Fragment
Estimate C-arm pose using FTRAC
Register C-arm images with CT volume
Display biomechanics & radiographic angles
Dewarp C-arm distortion
Before rectification
After rectification
Notes: Effects of dewarping  text, pelvis stretched (easy to see with iliac crest)

## Slide 28
Summary | Dependencies | Milestones | Deliverables | Experiment | Results
Imaging Subroutine
Segment BBs globally
Classify BBs as
FTRAC
Confidence
Fragment
Estimate C-arm pose using FTRAC
Register C-arm images with CT volume
Display biomechanics & radiographic angles
Dewarp C-arm distortion
At the moment, we use roipoly() to distinguish BB types
Notes: Global BB segmentation achieved using radial symmetry algorithm.
Had to adjust parameters to balance sensitivity and false positive rate.
Other ideas for BB classification  machine learning (SVMs, decision tree, Bayesian network), unique geometries, different materials

## Slide 29
Summary | Dependencies | Milestones | Deliverables | Experiment | Results
Imaging Subroutine
Segment BBs globally
Classify BBs as
FTRAC
Confidence
Fragment
Estimate C-arm pose using FTRAC
Register C-arm images with CT volume
Display biomechanics & radiographic angles
Dewarp C-arm distortion
At the moment, we use roipoly() to distinguish BB types
Notes: roipoly() not that bad a solution  doesn’t take long to do, and very accurate (see next slide)

## Slide 30
Summary | Dependencies | Milestones | Deliverables | Experiment | Results
Imaging Subroutine
Segment BBs globally
Classify BBs as
FTRAC
Confidence
Fragment
Estimate C-arm pose using FTRAC
Register C-arm images with CT volume
Display biomechanics & radiographic angles
Dewarp C-arm distortion
Results of BB classification
cyan = FTRAC
pink = confidence
blue = fragment
Notes: Minor bug to fix: how to do pose estimation when <9 FTRAC BBs are visible in x-ray image

## Slide 31
Summary | Dependencies | Milestones | Deliverables | Experiment | Results
Imaging Subroutine
Segment BBs globally
Classify BBs as
FTRAC
Confidence
Fragment
Estimate C-arm pose using FTRAC
Register C-arm images with CT volume
Display biomechanics & radiographic angles
Dewarp C-arm distortion
Existing software will perform 2D-3D registration using a correspondenceless expectation-maximization algorithm.
Notes: Screenshot = registration initialization GUI
  Need to initialize ECM algorithm
  Screenshot doesn’t show our data – femur from one our mentor’s past experiments. Just to show how 2d-3d registration works.

## Slide 32
Summary | Dependencies | Milestones | Deliverables | Experiment | Results
Imaging Subroutine
Segment BBs globally
Classify BBs as
FTRAC
Confidence
Fragment
Estimate C-arm pose using FTRAC
Register C-arm images with CT volume
Display biomechanics & radiographic angles
Dewarp C-arm distortion
We regard the registration from the optical tracker method as ground truth.
Notes: Notice rigid body on pelvis. Mention pivot calibration, touching DRB tip to confidence and fragment points.

## Slide 33
To Be Completed
Summary | Dependencies | Milestones | Deliverables | Experiment | Results
(1) BB Backprojection
Incorporate code to backproject BBs segmented in 2D x-ray images onto 3D CT volume
= projection of point in x-ray image
= point backprojected in 3D (world) space
We need to determine which BBs correspond to one another across multiple images…
Notes: Backprojection: we know BB locations in x-ray. This tells us where there are w.r.t. CT model.
Fragment movement: compute transformation in CT model space (reference frame derived from confidence BBs) from pre-op to post-op.

## Slide 34
To Be Completed
Summary | Dependencies | Milestones | Deliverables | Experiment | Results
(2) Locate Corresponding BBs
Determine which BBs correspond to one another across multiple pre-op (or post-op) image sets
Notes: Backprojection: we know BB locations in x-ray. This tells us where there are w.r.t. CT model.
Fragment movement: compute transformation in CT model space (reference frame derived from confidence BBs) from pre-op to post-op.

## Slide 35
To Be Completed
Summary | Dependencies | Milestones | Deliverables | Experiment | Results
(2) Locate Corresponding BBs
Image from: http://en.wikipedia.org/wiki/Epipolar_geometry
Corresponding BB in 2nd image will be on epipolar line
Epipolar Geometry
Locate one BB in 1st image
IMAGE 1
IMAGE 2
Notes: Backprojection: we know BB locations in x-ray. This tells us where there are w.r.t. CT model.
Fragment movement: compute transformation in CT model space (reference frame derived from confidence BBs) from pre-op to post-op.

## Slide 36
To Be Completed
Summary | Dependencies | Milestones | Deliverables | Experiment | Results
(3) Compute Fragment Movement
Determine fragment transformation between pre-op & post-op images
Transformation may be computed in model (3D) space or image (2D) space
(4) Better BB Classification
Create robust solution for differentiating FTRAC, confidence, & fragment BBs
Desired, but not necessary
Notes: Backprojection: we know BB locations in x-ray. This tells us where there are w.r.t. CT model.
Fragment movement: compute transformation in CT model space (reference frame derived from confidence BBs) from pre-op to post-op.

## Slide 37
References
1. Armand M, Lepisto J, Tallroth K, Elias J, Chao E. Outcome of periacetabular disease. Acta Ortho 2005;76(3):303-13.
2. Armiger RS, Armand M, Lepisto J, Minhas D, Tallroth K, Mears SC, Waites MD. Evaluation of a computerized measurement technique for joint alignment before and during periacetabular osteotomy. Computer Aided Surgery 2007;12(4):215-24.
3. Armiger RS, Armand M, Tallroth K, Lepisto J, Mears SC. Three-dimensional mechanical evaluation of joint contact pressure in 12 periacetabular osteotomy patients with 10-year follow-up. Acta Orthopaedica 2009 04;80(2):155-61.
4. Chintalapani G, Jain AK, Taylor RH. Statistical characterization of C-arm distortion with application to intra-operative distortion correction. Proc SPIE 2007;6509:65092Y.
5. David P, Dementhon D, Duraiswami R, Samet H. SoftPOSIT: Simultaneous pose and correspondence determination. International Journal of Computer Vision 2004 09/15;59(3):259-84.
6. Dementhon DF, Davis LS. Model-based object pose in 25 lines of code. International Journal of Computer Vision 1995;15(1):123-41.
7. Jain AK, Mustafa T, Zhou Y, Burdette C, Chirikjian GS, Fichtinger G. FTRAC---A robust fluoroscope tracking fiducial. Med Phys 2005 October 2005;32(10):3185-98.
8. Kang X, Taylor RH, Armand M, Otake Y, Yau WP, Cheung PYS, Hu Y. Correspondenceless 3D-2D registration based on expectation conditional maximization. Proc SPIE 2011 March 3, 2011;7964(1):79642Z.
9. Millis MB, Murphy SB. Osteotomies of the hip in the prevention and treatment of osteoarthritis. The Journal of Bone & Joint Surgery 1995;41:626-47.
10. Otake Y, Armand M, Armiger R, Kutzer M, Basafa E, Kazanzides P, Taylor R. Intraoperative image-based multi-view 2D/3D registration for image-guided orthopaedic surgery: Incorporation of fiducial-based C-arm tracking and GPU-acceleration. Medical Imaging, IEEE Transactions on 2011.