## Slide 1: Haptic Interface for Surgical Manipulator System Paper Presentation
Piyush Poddar
Mentors: Mike Kutzer, Ryan Murphy, and Mehran Armand
Computer Integrated Surgery II: Group 5
Spring 2012
Johns Hopkins University
Biomechanical and Image-guided Surgical Systems (BIGSS) Laboratory

## Slide 2
Kutzer, M. D. M., S.M. Segreti, C.Y. Brown, M. Armand, R.H. Taylor, and S.C. Mears. "Design of a New Cable-driven Manipulator with a Large Open Lumen: Preliminary Applications in the Minimally-invasive Removal of Osteolysis." International Conference on Robotics and Automation (2011): 2913-920. IEEEXplore. Web. 12 Apr. 2012. <http://ieeexplore.ieee.org/xpls/abs_all.jsp?arnumber=5980285&tag=1>.
* All unreferenced images are taken directly from this paper

## Slide 3
Problem:
An intuitive interface for a developed dexterous manipulator does not exist.
Solution:
develop such an intuitive interface using the PHANTOM® Premium while integrating haptic feedback

## Slide 4
PHANTOM Premium
Controller application
Snake robot & Z-Θ stage motors & sensors
Force feedback

## Slide 5: Paper Selection
Project Purpose
Physical Capabilities and Limitations
Methods for Modeling and Testing

## Slide 6
Purpose:
To describe the design, fabrication, control, and modeling of a manipulator designed for surgical removal of osteolysis.

## Slide 7: Background: Osteolysis
http://www.jbjs.org/data/Journals/JBJS/759/JBJA0791218340G03.jpeg
Osteolysis: Degeneration of bone tissue
Associated with wear particles following total hip replacement
May lead to bone fracture or implant loosening

## Slide 8: Background: Revision Surgery
http://ars.sciencedirect.com/content/image/1-s2.0-S0883540307003294-gr2.jpg
Lesion accessed through screw holes
Manually eliminate particle debris and osteolytic lesions
Surgically challenging: <50% of lesion typically grafted

## Slide 9: Goal of Procedure
Orient the lumen of the Dexterous Manipulator to different point in the lesion
Allow for introduction of surgical tools
  Metal Brushes
  Pincers
  Water Jets
  Suction

## Slide 10: Procedural Workflow
(1) Access screw hole by removing liner
(2) Insert rigid guide cannula
(3) Insert dexterous manipulator
(4) Insert and actuate tool tip through cannula
(5) Final cleaning and inspection
(6) Bone graft/Bone cement delivery

## Slide 11: Design Requirements
Compact
Open Lumen
Scalability
High-Dexterity
Large End-Effector Forces
Zimmer Trilogy Acetabular Shell with 6.60 mm diameter holes
http://www.zimmer.com/content/pdf/en-GB/Trilogy_Acetabular_System_Brochure_(97-6200-201-00_Rev_3)(07_2010).pdf

## Slide 12: Dexterous Manipulator Geometry
Outer Diameter = 5.99 mm (< 6.60 mm)
Notches force in-plane bending
  Simpler modeling and control
  Higher end-effector forces
2 drive cables housed in channels 180 apart to allow for marionette-like control

## Slide 13: Dexterous Manipulator Geometry

## Slide 14: Design Decision
How many notches to have and how much notch overlap?

## Slide 15: Manipulator Fabrication Process
Inner and Outer Nitinol Tubes Cut Using Lathe
Drive Cable Channels Cut on Outside of Inner Tube
Drive Cable Channels Cut on Inside of Outertube
Nesting of Tubes using Machine Vice
Notches cut using Electical Discharge Machining

## Slide 16: Manipulator Control
Stepper Motors Use for:
  Driving left and right cables
  Linear Actuation about Unbent Manipulator Axis
  Rotation Around Unbent Manipulator Axis

## Slide 17: System Sensors
Each cable connected to actuated via tension/compression load cell
CMOS FireWire Camera
  Manipulator stays within field-of-view
  Question: How to define zero position of each of 4 actuators?

## Slide 18: Z-Calibration
2 Limit Switches on both ends of Translating Component
Steps:
  Actuate to each Limit Switch
  Choose Fixed Intermediate Point as Zero Position

## Slide 19: -Calibration
Steps:
Bend manipulator
Rotate by about 360
  Take image at each step
  Segment image (using Intensity-Based Segmentation)
    Measure Distance from DM Base to End-Most Pixel for left and right side
Mark 0 Position at Bending Plane Orthogonal to Camera Axis

## Slide 20: Drive-Cable Length Calibration
Each Drive Cable has 4.4N of Counter-Tension
  Total Bend Angle Estimated
  Slack cable on bend-ward side by # of steps proportional to bend angle
  Actuate opposite cable to maintain 4.4N Counter-Tension
  Cable on bend-ward side set to 4.4N counter tension and angle estimated again
  Repeat until angle is about 0with Counter-Tension in each cable at 4.4N

## Slide 21: Kinematic Modeling
Purpose: Motion Planning and Future Design Optimization
Modeled as a Series of Pin Joints that connect a sequence of Rigid Vertebrae

## Slide 22: Experimentation and Testing
Experiment 1: Investigate the Accuracy of the Kinematic Model
Methodology:
Acquire 120 images of manipulator configurations
Acquire aforementioned 58 points
For first 100 mages, use Linear Least-Squares to force equality of constant parameters () and calculate .
Find error between selected point and calculated point

## Slide 23: Experiment 1: Verifying the Kinematic Model
Using First 100 training images
Using determined parameters on next 20 test images

## Slide 24: Experimentation and Testing
Experiment 2: Define Relationship Between End-Effector Forces and Cable -Tension
Methodology:
At bend configurations of :
Apply force gauge in Eastward, Westward, and Southward directions
Bring gauge in contact with end-effector while maintaining 0.00 N reading
Step motor and record load cell and force-gauge data
Perform 3 trials with 30 data points each

## Slide 25
Experiment 2: End-Effector Force and Cable-Tension Relationship

## Slide 26
Experiment 2: End-Effector Force and Cable-Tension Relationship

## Slide 27
Experiment 2: End-Effector Force and Cable-Tension Relationship

## Slide 28
Experiment 2: End-Effector Force and Cable-Tension Relationship

## Slide 29: Out of the Bend-Plane
Similar Test
  Only at
  rotated by 90 to bring bend-plane same axis as camera

## Slide 30: Conclusions
Design objectives were met
No signs of fatigue or failure
Promising kinematic results
Some variability between identical trials
Path-planning work using kinematic model shows 85%-95% coverage rates

## Slide 31: Next Steps
Optimization of Manipulator Geometry
3D bending via addition of more drive cables
Predicting cable lengths from known-position and vice-versa
Develop intuitive, haptic control interface!

## Slide 32: Relevance and Importance to Project
Learned about Cable-Properties of Manipulator
Range of Manipulator Motion
Possible Kinematic Model
Calibration Steps

## Slide 33: My Critique: The Good
Wide breadth of information
Relatively Large Sample Size during Experiments
Good Use of Figures and Plots
Clear and Concise

## Slide 34: My Critique: The Less Good
Some important features were not explained
  Image-based segmentation
  Current control interface
  Path-planning work indicating up to 95% coverage
Sometimes difficult to understand
  Kinematic Model
  Drive-Cable Length Calibration

## Slide 35: All in All
Very useful in developing background in project
Helped to realize overall purpose
What considerations have gone into the design
Possible modeling and testing ideas

## Slide 36: Questions or Comments?
Thank You!