## Sheet1
| Severity Scale: |  |  | Occurrence scale: |  |  | Detection scale:
| 1 | No harm to patient or operator and minimal disturbance to procedure |  | 1 | Extremely unlikely (should not occur during trials) |  | 1 | Detection certain (fault will always be detected)
| 2 | Minimal harm/chance of harm to patient or operator (very minor injury or disturbance to procedure) |  | 2 | Minimal chance of occurrence (may occur once during whole set of trials) |  | 2 | Detection probable (fault is likely to be detected)
| 3 | Moderate harm/ chance of harm to pateint or operator (minor injury or disturbance to procedure) |  | 3 | Moderate chance of occurrence (may occur once in 10 procedures) |  | 3 | Detection possible (fault has approximatley 50% chance of being detected)
| 4 | Serious harm/chance of harm to patient or operator (serious injury or significant disturbance of procedure) |  | 4 | Likely chance of occurrence (may occur once in every 2-3 procedures) |  | 4 | Detection unlikely (fault will probably be undetected)
| 5 | Severe harm/chance of harm to patient or operator (life threatening or procedure failure) |  | 5 | Certain occurrence for each procedure |  | 5 | Detection impossible (fault will never be detected)
Item/Function | Potential Failure Mode | Potential Effects of Failure | S | Potential Cause(s) | How Failure is Detected | O | Current Controls | D (detection r | RPN (risk priority number) | Reccomended Actions
Robot | Water gets into motoc controller enclosure and causes damage to controller | Controller stops working | 2 | Cover of controller is incorrectly installed and someone spills water on cotroller | Watchdog timer on the PC will time out when controller is unresponsive and shut down the system via the USB relay | 1 | Water resistant enclosure | 1 | 2 | Training and proper maintainence of seals
| Limit switch mechanically fails during operation | System runs into software limit and is stopped | 1 | Limit switch fatigue, manufacture defect | If the soft limit fails, this will be detected as a movement error by the motor controller, otherwise failure will be undetected. | 2 | Soft limits act as backup | 5 | 10 | Stop and send robot for maintainence
| Pot has electrical or mechaincal failure before operation begins | Would not be able to cross check the encoders | 2 | Pot fatigue or manufacture defect | Detect by comparing to expected pot values stored in program | 2 | Cross check with encoder during calibration procedure (can't be detected during operation) | 1 | 4 | Stop and send robot for maintainence
| Pot has electrical or mechaincal failure during operation | Would not be able to cross check the encoders | 2 | Pot fatigue or manufacture defect | Would show a cross checking error | 2 | Cross check with encoder during operation | 1 | 4 | Stop and send robot for maintainence
| Encoder has electrical or mechanical failure before operation begins | Motors would not be able to function correctly | 2 | Encoder fatigue or manufacture defect | Detected during calibration procedure | 2 | Do calibration procedure and cross check with pots | 1 | 4 | Stop and send robot for maintainence
| Encoder has electrical or mechanical failure during operation | Motors would not be able to function correctly | 2 | Encoder fatigue or manufacture defect | Detected by cross checking with pots | 2 | Cross checking with pots | 1 | 4 | Stop and send robot for maintainence
| Both encoder and pot have electrical or mechanical failure | Motors would not be able to function correctly | 3 | Encoder or pot fatigue or manufacture defect | Detected by the motor controller motion error | 1 | Motor controller motion error | 2 | 6 | Stop and send robot for maintainence
| Motor failure due to broken wire or short | Motor would stop moving | 2 | Motor manufacture defect | Detected by motor controller motion error | 2 | Motor controller motion error | 1 | 4 | Stop and send robot for maintainence
| Mechanical break between the motor and joint | Robot could move incorrectly | 3 | Mechanical fatigue, pushed beyond load capacity | Detected by cross checking with pots | 2 | Cross checking with pots | 1 | 6 | Stop and send robot for maintainence
| Arm joint becomes loose | Arm could fall and injure patient or operator | 3 | Gross mishandling of the robot | Robot falls during procedure | 1 | Follow handling instructions, Robot is positioned so robot mass is not over patient | 1 | 3 | Stop and send robot for maintainence
| Object blocking arm or internal jam in arm | Could cause damage to robot | 3 | Obstacles in range of motion, internal mechanical problem | Motor controller motion error | 2 | Follow setup instructions | 1 | 6 | Stop and send robot for maintainence
| Broken wire, short causes robot to become electrified | Weak electric shock (12 V) | 2 | Broken wire | Fuse would blow | 1 | chassis ground wire;only 12 V;fuse guards against large short; scope electrically insulated | 1 | 2 | Proper maintainence of wires
| Limit switch mechanically fails before operation begins | Robot would run into end of limit during startup procedure | 2 | Limit switch fatigue, manufacture defect | Motor controller detects that there is a motion error durin startup, aborts startup | 2 | Startup calibration procedure will detect | 1 | 4 | Send robot for maintainence
Computer Failure | Ethernet failure: motor controller loses communication with PC | Robot would stop responding to commands | 3 | Ethernet cord is disconnected | Watchdog timer | 2 | PC watchdog timer times out and cuts power to motors | 1 | 6 | stop and check ethernet cord, maintainence if nothing appears wrong with cord
| Computer crash | Computer stops responding to user commands, cannot command robot | 3 | Software errors | Motoc controller watchdog timer | 2 | Motor controller watchdog timer times out and cuts power to motors | 1 | 6 | Stop, restart and reload calibration tables
| Power surge causes damage to computer | No effect or computer failure | 3 | Power surge in building | Computer stops working; fuse blows | 1 | isolated power supply; fuse will catch | 1 | 3 | Use a surge protector power strip
Installation Error | Incorrect Attachment to bedrail | Could injure patient or surgeon | 4 | Incorrect installation | Check stability in setup procedure | 1 | Check robot stability during setup; large margin for error in bedrail attachment system | 1 | 4 | Correct training for setup and surgeon
| Joints not locked correctly | Scope pulled from patients mouth; damage scope | 2 | Incorrect installation | Check stability in setup procedure | 2 | Check robot stability during setup; all joints that could move due to gravity have friction collars to prevent unintended motion | 1 | 4 | Correct training for setup and surgeon
| Scope cord not slack enough and becomes caught on something | Could damage scope | 3 | Incorrect installation | Check range of motion of robot before procedure | 2 | Proper setup | 1 | 6 | Correct training for setup and surgeon
| Friction collars not tightened correctly | Damage scope, robot; injure patient; injure operator | 3 | Incorrect maintainence | Check friction collars during setup | 1 | Proper maintainence and checks during setup | 1 | 3 | Correct maintainence and correct training for setup
| Unsanitary draping/undraping procedure | Spread of contamination between patients | 4 | Incorrect procedure | Patient contracts infection | 1 | Draping and cleaning procedure; disposable parts | 2 | 8 | Follow proper draping and cleaning procedures; secondary cleaning; double draping?
Maintainance | Seals on robot degrade | cleaning fluid gets into the robot during cleaning | 2 | Material fatigue | Visually observed during maintainence | 1 | Maintainence and inspection | 1 | 2 | Proper maintainence
| Friction collars degrade | Robot comes loose during adjustment | 3 | Material fatigue | Inspection and testing during maintainence | 2 | Maintainence and inspection | 1 | 6 | Follow maintainence procedure
Joystick Failure | Joystick mechanical failure: joystick cannot change switch settings | User can't control the robot | 2 | Material fatigue or mishandling | User detects robot's unresponsiveness | 1 | Maintainence and inspection/check in startup procedure | 1 | 2 | Stop operation and send robot for maintainence
| Joystick double switch failure | incorrect command | 3 | short circuit, broken wire, mechanical failure | Robot not responding properly to surgeon commands | 1 | Emergency stop button, checks for inconsistent joystick commands | 2 | 6 | Emergency stop, send for maintainence