# Summary for Apr. 26 US calibration

Improvements made in comparison with the previous version:

Settings:

- Change settings of US machine: freq 6.6, persist 0, dynamic range less, focus was at each time put at the pointer point, depth 9cm, stored b8 image.

- Used a sharp pointer

- setup was made more stable by using a table and passive arms:

4. reprinted markers (should not be taped)

5. The pointer was kept by a passive arm to ensure synchronization.

## Analysis:

Considered speed of sound compensation, cut images and measured dimension using US image. Checked if the rotation matrix is RzRyRx which is correct.

## Resutls:

The standard deviation of probe position was:

0.2384    0.2539    1.0596

In millimeters. The average of probe position and angles are:

-45.5117  -97.0130  553.9648   -0.0857   -0.4650   -2.9511

Bad data:

2, 53 : not 10 frames of US images

47: arrays on top prohibit auto segmentation

42: no pointer data

Transformations:

frame1 =

0.2844    0.8520   -0.4396  142.2242

-0.2335   -0.3832   -0.8937 -102.1007

-0.9298    0.3568    0.0899   87.6718

0         0         0    1.0000

frame2 =

0.2844   -0.2335   -0.9298   17.2292

0.8520   -0.3832    0.3568 -191.5770

-0.4396   -0.8937    0.0899  -36.6060

0         0         0    1.0000

Frame2 is the transformation from marker coordinate system to the us image coordinate system. Now looking at the following image we will see the two transformations make sense except that I still don’t understand why frame2(1,4) = 17.22 and frame1(3,4)=87.67. In my opinion they should not be that different.

My other observation is that the reason for having a big error along x-axis in version 2 calibration (which most probably will be observed in v3 as well) is the image thickness! Because I could see that the pointer_tip_x in camera coordinates is varying from -30 to 30 mm. That can be the reason for the mismatch in the above transformations too.

## Verification of data v3, data collection 3/5/2013 analysis 5/5/2013

6-12: collected with known cross point position in tracker

1-5 (second set): collected without knowing position of point and I did not like the data it seemed to be not accurate.

Results of evaluation: (first set 6-12):

Precision

stderror =

4.3467    1.2044    1.3065

Mean_cloud=

37.3689   86.5109  479.2239

Accuracy:

pointer =

32.9422   90.9091  471.5404

mean_cloud =

37.3689   86.5109  479.2239

mean_error =

-4.4267    4.3982   -7.6835