## Slide 1: 3D Scanning Applications for Intraoperative Cranioplasty Implant Modification
Alex Mathews ∙ Joshua You
Computer Integrated Surgery 2015
Mentors: Dr. Armand, Dr. Gordon, Dr. Murphy, Captain Grant

## Slide 2: Team Members
Alex Mathews
Joshua You
Mentors
  Dr. Armand – JHU APL
  Dr. Gordon – JHMI
  Dr. Murphy – JHU APL
  Captain Grant – Walter Reed

## Slide 3: Overview
Project Summary
Progress
  Dependencies
  Milestones
Deliverables
Timeline
Upcoming Milestones

## Slide 4: Project Summary
Implementing the iPad mounted Structure Sensor to intraoperatively scan a patient’s head during a cranioplasty procedure
Key goals:
  Reduce operation time
  Improve accuracy and fit of CCI
Key topics:
  3D Scanning
  3D Mesh Segmentation
Notes: CCI = Customized Cranial Impant

## Slide 5: Progress
Dependencies, Milestones

## Slide 6: Dependencies
Structure Sensor - Completed ✓
  Purchased by Dr. Armand
Mock Operations - Completed ✓
  Once a month with Dr. Gordon
Test Objects - Completed ✓
  Calibration objects and Phantom Skull
  Printed at the DMC and by Dr. Ryan Murphy
iPad - Completed ✓
  Personal iPad

## Slide 7: Milestones
Date | Milestone | Status
Feb 28 | Color Segmentation algorithm completed | ✓
Mar 14 | .obj area of interest algorithm completed | ✓
Mar 28 | Phantom skull testing | ✓
Apr 4 | Mock Operation: Test scans in surgical environment (Completed March 7) | ✓
Apr 2 | Checkpoint presentation | ––
Apr 18 | Complete Procedure Integration | ––
May 2 | Mock Operation: Test integration modules | ––

## Slide 8: ✓ Scanning and Color Segmentation
Scanner is used in uplink mode
  Skanect
  Implements Structure Sensor
File is imported into modules using the .obj format
  ASCII file type
  Vertices with RBG color data
Colors are segmented during import of .obj file

## Slide 9: ✓ Phantom Skull Testing
Scanner Output
*Note colored dots placed on skull
Ground Truth Skull
*Defect colored to provide contrast in this presentation
Notes: Ground truth skull was generated from compilation of CT data

## Slide 10: –– Integration – Surgical Workflow
* Only necessary if required by CT surface model registration
** Requires Preoperative 3D Scan if facial structures hidden due to patient positioning
Preoperative CT Scan
Before Surgery
CT Surface Model Built
Implant Designed
Implant Fabricated
Surgery
Skull Fragment Removed
Intraoperative 3D Scan
Surgery Prep
Head Shaved
Preoperative 3D Scan*
Reflective Markers
Registration to CT Surface Model**
Defect CT Coordinates
Identified
Crest Identification
Intraoperative Processing

## Slide 11: –– Integration – Registration
Scan to CT Surface Model Registration
  Implant is designed in CT coordinates
  Registration required to determine defect coordinates in the CT space
  Will use bony structures of the face to determine gross registration
Use ICP for refinement of transformations

## Slide 12: –– Integration – Registration
Scan to Scan Registration
  Required if the face is not in view due to patient positioning
  Use colored dots for gross registration ✓, ICP for refinement ––
Gross registration of isolated dots from two separate scans

## Slide 13: –– Integration – Registration (cont.)
Next steps:
  Use ICP to refine registration for Scan to Scan
  Identify optimal bony structures for Scan to CT Surface Model registration
  Use ICP for Scan to CT Surface Model registration
  Looking into more radiant colored dots for segmentation
Validation steps:
  Need to provide data on accuracy of scan and registration
  Currently unknown, but should be within millimeter accuracy
  Need to determine variation between scans

## Slide 14: –– Integration – Crest Identification
Past Attempts - Using previously developed programs (Mayavi)
  Issues
    Still unsure if its possible to extract contour information
    Orientation dependent
Past Attempts – Extracting contour lines using ring relationship based method
  Issues
    Threshold difficult to determine
    Orientation dependent

## Slide 15: –– Integration – Crest Identification (cont.)
Future Method #1
  Use centroid of the mesh to determine the center of the defect
  Use a collection of surrounding points to determine a normal vector to surface
  Calculate normal vectors of all other points using surrounding points
  Use angle of normal vectors to determine walls of defect
Worst Case Scenario – Manual Input
  Surgical assistant manually selects points on the work station
  Equivalent to digitizing defect using a pointer in the current procedure

## Slide 16: Deliverables
Updates to project goals

## Slide 17: Previous Deliverables
Minimum
Segmentation
  Color
  Depth
Mesh focusing algorithm through crest detection
Expected
Ability to use the scanner in uplink mode to workstation
Integration with laser projection modules
Maximum
Testing in the Mock Operation Room
Tablet application with 3D scanning and processing

## Slide 18: New Deliverables
Minimum
Segmentation
  Color ✓
  Depth ✓
Mesh focusing algorithm through crest detection ––
Expected
Ability to use the scanner in uplink mode to workstation ✓
Integration with laser projection modules ––
Maximum
Testing in the Mock OR ✓––
Tablet workstation application with 3D scanning and processing

## Slide 19: Timeline
Updates to project timeline

## Slide 20: Previous Timeline
 | February |  |  | March |  |  |  | April |  |  |  | May
Segmentation Algorithm Development |  |  |  |  |  |  |  |  |  |  |  | 
Working with .obj file, decide on data structures for import |  |  |  |  |  |  |  |  |  |  |  | 
Color identification algorithm |  |  |  |  |  |  |  |  |  |  |  | 
Refocusing .obj file algorithm |  |  |  |  |  |  |  |  |  |  |  | 
Defect edge identification using depth algorithm |  |  |  |  |  |  |  |  |  |  |  | 
Bottom edge, top edge joining and mesh output |  |  |  |  |  |  |  |  |  |  |  | 
Imaging Development |  |  |  |  |  |  |  |  |  |  |  | 
Define procedure for obtaining scan with skannect |  |  |  |  |  |  |  |  |  |  |  | 
Build test object with defect |  |  |  |  |  |  |  |  |  |  |  | 
Use test object for scans and determine accuracy |  |  |  |  |  |  |  |  |  |  |  | 
Test algorithm on object |  |  |  |  |  |  |  |  |  |  |  | 
Test algorithm on phantom skull |  |  |  |  |  |  |  |  |  |  |  | 
Test at Mock Operation |  |  |  |  |  |  |  |  |  |  |  | 
Procedure Integration |  |  |  |  |  |  |  |  |  |  |  | 
Develop necessary modules for scanner communication |  |  |  |  |  |  |  |  |  |  |  | 
Test at Mock Operation |  |  |  |  |  |  |  |  |  |  |  | 
Application Development |  |  |  |  |  |  |  |  |  |  |  | 
SDK, OpenNI exploration |  |  |  |  |  |  |  |  |  |  |  | 
Prototype Application Development |  |  |  |  |  |  |  |  |  |  |  | 

## Slide 21: New Timeline
 | February |  |  | March |  |  |  | April |  |  |  | May
Segmentation Algorithm Development |  |  |  |  |  |  |  |  |  |  |  | 
Working with .obj file, decide on data structures for import |  |  |  |  |  |  |  |  |  |  |  | 
Color identification algorithm |  |  |  |  |  |  |  |  |  |  |  | 
Refocusing .obj file algorithm |  |  |  |  |  |  |  |  |  |  |  | 
Defect edge identification using depth algorithm |  |  |  |  |  |  |  |  |  |  |  | 
Bottom edge, top edge joining and mesh output |  |  |  |  |  |  |  |  |  |  |  | 
Imaging Development |  |  |  |  |  |  |  |  |  |  |  | 
Define procedure for obtaining scan with skannect |  |  |  |  |  |  |  |  |  |  |  | 
Build test object with defect |  |  |  |  |  |  |  |  |  |  |  | 
Use test object for scans and determine accuracy |  |  |  |  |  |  |  |  |  |  |  | 
Test algorithm on object |  |  |  |  |  |  |  |  |  |  |  | 
Test algorithm on phantom skull |  |  |  |  |  |  |  |  |  |  |  | 
Test at Mock Operation |  |  |  |  |  |  |  |  |  |  |  | 
Procedure Integration |  |  |  |  |  |  |  |  |  |  |  | 
Develop necessary modules for scanner communication |  |  |  |  |  |  |  |  |  |  |  | 
Test at Mock Operation |  |  |  |  |  |  |  |  |  |  |  | 
Application Development (Workstation) |  |  |  |  |  |  |  |  |  |  |  | 
SDK, OpenNI Skannect exploration |  |  |  |  |  |  |  |  |  |  |  | 
Prototype Application Development |  |  |  |  |  |  |  |  |  |  |  | 

## Slide 22: Upcoming Milestones
Completion of defect isolation using crest identification algorithm
Implementation of ICP for Scan to Scan registration
  Testing and validation
Implementation of ICP for Scan to CT Surface Model registration
  Testing and validation against ground truth skull