## Slide 1: DVRK stereo camera calibration and model registration -Checkpoint Presentation-
Group 12: Peter Ahn and Mengze Xu Mentors: Preetham Chalasani and Anton Deguet April 6, 2017

## Slide 2: Background
Goal: Register surfaces to the robot (Patient Side Manipulator, PSM)
Current Method: Move robot to touch the surface
What we want to do: Using calibrated stereo camera to substitute touching
Notes: Surfaces such as the skin or organ surfaces of the patient

## Slide 3: Background
Camera
PSM
Phantom
Complete Hand-Eye Calibration
Detect and register phantom surface to PSM
Desired accuracy in several millimeters
Camera
PSM
Phantom
Previous Camera Mount
Notes: Surfaces such as the skin or organ surfaces of the patient

## Slide 4: Current Deliverables
Minimum
Hand-Eye Calibration and Accuracy Test
[Completed on March 31]
Expected
Predefined Surface Detection and Accuracy Test
[In Progress]
Maximum
Unknown Surface Detection and Accuracy Test
[Not Started]

## Slide 5: Progress – Hand Eye Calibration
A: Marker pose to camera frame (ArUco package)
X: Marker pose to end effector
B: End effector pose to PSM base (DVRK API)
Z: Camera Pose to PSM Base

## Slide 6: Progress – Hand Eye Calibration Setup

## Slide 7: Progress – Data Collection Process
In practice, e is a key parameter
Discard Data

## Slide 8: Calibration Consistency
Consistency of Same X and Y |  |  |  | 
 | RX/rad | tX/mm | RY/rad | tY/mm
e = 0.01 | 0.042 | 20.4 | 0.048 | 19.6
e = 0.001 | 0.043 | 3.5 | 0.023 | 2.7
Result of e = 0.001 is not bad, but it’s very hard to collect data
e = 0.003 is used later
Convenience vs. Better Accuracy
Notes: It may seem like we have just changed the wording of the expected deliverables, but we will explain what we changed during completing our deliverables

## Slide 9: Calibration Consistency
Consistency of Same X with different Y |  | 
Error | RX/rad | tX/mm
e = 0.003 | 0.054 | 10.5
Consistency of Same Y with different X |  | 
Error | RY/rad | tY/mm
e = 0.003 | 0.025 | 3.1
Notes: It may seem like we have just changed the wording of the expected deliverables, but we will explain what we changed during completing our deliverables

## Slide 10: Calibration Consistency
For the case of same Y, we move the adapter (X) for about 12mm and the change of tX is 12.7 mm
Notes: It may seem like we have just changed the wording of the expected deliverables, but we will explain what we changed during completing our deliverables

## Slide 11: Calibration Accuracy
error = [-0.4, 2.8, -1.3] mm
|| error || = 3.1 mm
Notes: It may seem like we have just changed the wording of the expected deliverables, but we will explain what we changed during completing our deliverables

## Slide 12: Phantom Detection Accuracy (via Touching)
error = [34.6, 55.7, -13.7] mm
|| error || = 67 mm
Mainly due to the poor result of depth information
Notes: It may seem like we have just changed the wording of the expected deliverables, but we will explain what we changed during completing our deliverables

## Slide 13: Result of Edge Detection
Notes: It may seem like we have just changed the wording of the expected deliverables, but we will explain what we changed during completing our deliverables

## Slide 14: Result of Edge Detection
RGB Image
Notes: It may seem like we have just changed the wording of the expected deliverables, but we will explain what we changed during completing our deliverables

## Slide 15: Result of Edge Detection
HSV Image (color threshold to eliminate background)
Notes: It may seem like we have just changed the wording of the expected deliverables, but we will explain what we changed during completing our deliverables

## Slide 16: Result of Edge Detection
Converted back to RGB Image
Notes: It may seem like we have just changed the wording of the expected deliverables, but we will explain what we changed during completing our deliverables

## Slide 17: Result of Edge Detection
Sectioned Binary Image
Notes: It may seem like we have just changed the wording of the expected deliverables, but we will explain what we changed during completing our deliverables

## Slide 18: Result of Edge Detection
Edge Image (Sobel Filter)
Notes: It may seem like we have just changed the wording of the expected deliverables, but we will explain what we changed during completing our deliverables

## Slide 19: Result of Edge Detection
Line Image (Hough Transform)
Notes: It may seem like we have just changed the wording of the expected deliverables, but we will explain what we changed during completing our deliverables

## Slide 20: Technical Approach Changed
Expected Deliverable
Detect the edge of phantom with single stereo camera and register with model
Set up two R200 RealSense camera to build a stereo camera system and detect edge points
Maximum Deliverable
Find correspond surface points in two camera
Notes: It may seem like we have just changed the wording of the expected deliverables, but we will explain what we changed during completing our deliverables

## Slide 21: Future Work
Set up two calibrated camera systems
Complete ICP algorithm for detected edge points and test for accuracy
Complete algorithm to find correspond points to estimate point cloud of unknown surface
Notes: It may seem like we have just changed the wording of the expected deliverables, but we will explain what we changed during completing our deliverables

## Slide 22: Updated Dependencies
Dependencies |  | How to Resolve | Status
Hardware | dVRK System | Access to LCSR Lab | Resolved
 | Stereo / Depth Camera | Access to LCSR Lab | Resolved
 | Two R200 Realsense Cameras | Find a way to connect two stereo cameras into one USB hub | Pending
 | 3D Printer | Access to 3D Printer Room | Resolved (Peter)
 | Opaque Phantom | Develop another phantom with different coloring dye | Pending [Not immediate concern]
Software | ROS | Open source codes | Resolved
 | ArUco / OpenCV Package | Open source codes | Resolved
 | Solidworks (CAD) | WSE IT website | Resolved (Peter)
 | MATLAB | WSE IT website | Resolved

## Slide 23: Updated Timeline
Plan Date | Progress | Actual Complete Date
2/18 | Finish proposal and presentation | 2/18
2/25 | Prepare hardware and environment | 3/05
3/11 | Complete codes for hand-eye calibration | 3/14
3/18 | Test codes and estimate accuracy | 3/31
4/01 | Complete contour detection | 3/25
4/08 | Set up 2 camera system and calibrate | 
4/11 | Test codes and estimate accuracy for expected deliverables | 
4/15 | Complete codes for finding correspond points | 
4/22 | Test codes and estimate accuracy for expected deliverables | 
 | Leave time for other methods | 

## Slide 24: Original Timeline

## Slide 25: Updated Timeline

## Slide 26: Updated Timeline