## Slide 1: Force Control Algorithms for Sclera Eye Surgery
Team Members: Ankur Gupta, Saurabh Singh
Mentors: Dr. Iulian Iordachita, Dr. Marin Kobilarov and Dr. Russell H. Taylor
Mar 28, 2017
Group #20

## Slide 2: Background
Taken in part from
Xingchi He
Dissertation Defense
June 30th, 2015

## Slide 3: Background: Admittance Control
Constant Admittance Control
  Isotropic admittance
  Haptic feedback using force scaling
  The robot cooperative control law incorporates scaled sclera force measurement
Fss
Fhs

## Slide 4: Background: Variable Admittance Control
Anisotropic admittance to render virtual fixture
6-DOF motion
4-DOF motion
RCM

## Slide 5
Based on the surgeon’s behavior i.e. application of force on sclera, devise a function from the force scaling mode of the co-operative control to the full RCM mode.
Currently, the surgeon based on his experience cancels out the possibility of translational motion after reaching certain depth in the eye .
Therefore, it is proposed to devise experiments which will capture the behavior of several surgeons to learn about their surgical procedure.
Based on that collected data and according to current depth of tool from sclera restrict the surgeon from making movements which are detrimental to sclera .
Overview of the Objective
From the data collected, learn
  Transition function
  Lower and Upper depth bounds after sclera where the operating modes switch
  β , the parameter which reduces degree of freedom

## Slide 6
The transition function between the lower bound and the upper bound depth of the sclera depth is considered to be the straight line function but a study has to be conducted for the nonlinear functions as it is assumed that it won't be a linear function for human operative conditions.
Overview of the Objective
From the data collected, learn
  Transition function
  Lower and Upper depth bounds after sclera where the operating modes switch
  β , the parameter which reduces degree of freedom
7
6-DOF motion
4-DOF motion
RCM

## Slide 7
Design of Experiments
1) Tool in normal position: Tool to be placed in normal position(zero pitch and yaw) and to be moved 1 mm to take 10 readings . The initial tool insertions depth is 10 mm. This experiment will reflect the behavior of the force with insertion depth.
2) Tool inclined at an angle and brought to normal: Tool to be placed at an angle(non zero pitch and roll). This experiment will help us to see the behavior of the force in making the marked point normal. This behavior will reflect the choice of the surgeon to use depth, force and the angle of the tool to rotate the eyeball.
Technical Approach: Design of Experiments

## Slide 9
Tool at normal position in the eye i.e., pitch and roll of the end effector are zero

## Slide 10
Tool with non zero pitch and roll

## Slide 11
Development of Eyeball of phantom
Collection of Force data on Sclera from surgeons will be based on an experiment to maintain repeatability and same data generation environment
Fours inside the eyeball of the phantom are marked with different colours around the center point of retina
Surgeons will be apply forces to rotate the eyeball from one coloured marker to the centre point of the retina
  Experimentation Setup

## Slide 12
Reduction in friction while motion of eyeball in phantom eye socket
Ultrasound Gel is being used as lubricating agent
Other available lubricants in AMIRO lab. were water based hence Ultrasound Gel
Water based lubricants loose their properties easily
Procured measuring cup to measure gel and maintain same experimental conditions in future
Experimentation Setup

## Slide 13
Force Calibration Validations
Force data depends on calibration of the tool.
Thus, validating the calibration of the tool is important
This done by placing the tool on measuring scale and applying forces on tool shaft
The calibration of tool is acceptable if force values from tool and scale are within acceptable bounds.
Experimentation Setup

## Slide 14
Distance from Sclera to Tip Validation
Checked the consistency of the software to detect the position of the sclera in the end-effector frame
Placed a ruler and changed the point of force application on several instances
The results showed that the change in point of force application reflected as change in sclera position
Experimentation Setup

## Slide 15
Experimentation Setup

## Slide 16
Restored the system back to its running state
Made ourselves familiar with code and the robot
Reproduced the pivot point calibration experiment mentioned in the Xinchi He thesis to check functionality of the Sclera (Force Admittance) Mode.
Ran tests to find any offsets with in motion of the robot. It was found that the robot has an offset of 8 degrees in roll DOF in its homing position. Efforts to remove those offsets are in progress
Overview of Progress

## Slide 17: Detailed Plan

## Slide 18
Original Project Timeline

## Slide 19
Original Project Timeline
2017
Feb
Mar
Apr
May
Plan Presentation
28/2/2017
Checkpoint Presentation
28/3/2017
Project final report
18/5/2017
Knowledge Transfer
15/3/2017
System Setup
8/4/2017
Proof of Concept
2/5/2017
Robust Validation
18/5/2017
Write Paper
30/5/2017
Procure source code
20/2/2017 - 27/2/2017
Contact XingChi He
20/2/2017 - 27/2/2017
Procure Hardware, Microscope, FBG, Fiber Optic sensor
20/2/2017 - 14/3/2017
Calibrate Fiber-Optic Force sensing tool
14/3/2017 - 23/3/2017
Identify source code pointing to Sclera force sensing
27/2/2017 - 14/3/2017
Restore the full setup to working state
14/3/2017 - 8/4/2017
14/3/2017 - 22/4/2017
Port backbone code to C++
Introduce non-linear behaviour to force feedback
23/3/2017 - 24/4/2017
Proof of concept with ‘untrained’ hands
24/4/2017 - 2/5/2017
Data collection from doctors
2/5/2017 - 10/5/2017
Validation of approach with new this data
11/5/2017 - 18/5/2017
24/4/2017 - 26/5/2017
Write Research Paper
Today

## Slide 20
Revised Project Timeline: Significant Changes Made to Original
2017
Today
Feb
Mar
Apr
May
Plan Presentation
28/2/2017
Project final report
18/5/2017
Knowledge Transfer
15/3/2017
System Setup
8/4/2017
Proof of Concept
2/5/2017
Robust Validation
18/5/2017
Write Paper
30/5/2017
Procure source code
20/2/2017 - 27/2/2017
Contact XingChi He
20/2/2017 - 27/2/2017
Procure Hardware, FBG, Fiber Optic sensor
20/2/2017 - 8/3/2017
Calibrate Fiber-Optic Force sensing tool
14/3/2017
Identify source code pointing to Sclera force sensing
27/2/2017 - 21/3/2017
Restore the full setup to working state
14/3/2017 - 24/3/2017
Port backbone code to C++
14/3/2017
Introduce non-linear behaviour to force
23/3/2017 - 18/4/2017
Proof of concept with ‘untrained’ hands
18/4/2017 - 28/4/2017
Data collection from doctor
28/4/2017 - 1/5/2017
Validation of approach with new data
1/5/2017 - 6/5/2017
Report and Poster Prep
6/5/2017 - 17/5/2017
24/4/2017 - 26/5/2017
Write Research Paper

## Slide 21: Deliverables
Minimum:
Procured software and Algorithm bridge source code
Ported backbone code to C++, if required
Procured and Calibrated Fiber-Optic Force sensing tool (Replace Fiber Optic tool, with low cost workable tool, if FO not procured)
Identified and documented source code that pertains to Sclera force sensing, drivers, APIs, and backbone algorithm
Restored the full setup to state, working with all tools and sensors attached.
Analysis of force feedback curve with multiple and unique coefficients
Sclera Distance and Force Validation Data Analysis. [Added]
Green: Completed
Yellow: In Progress
Blue: Not Started
Black: Unlikely

## Slide 22
Expected:
Introduction of non-linear behavior to force feedback
Proof of concept with ‘untrained’ hands on Chorioallantoic membrane or synthetic sclera-impersonator membrane, such that forces seem natural in the process.
Green: Completed
Yellow: In Progress
Blue: Not Started
Black: Unlikely
Deliverables

## Slide 23
Maximum:
Data collection from ‘experts’ or doctors
Validation of model with ground truth force feedbacks.
Deliverables

## Slide 24: Dependencies(Almost there)
No. | Dependency | Resolve by | Status | Plan B
1 | Develop complete understanding of problem and survey for possible approaches, based on feedback from mentors. | 3/3/2017 | Resolved | -
2 | Procure codebase | 02/24/2017 | Resolved | -
4 | Xingchi’s appointment to discuss inconsistencies in impl. and thesis | 3/31/2017 | Resolved | -
5 | Fiber Optic sensor from Dr. Iordachita. | 3/3/2017 | Resolved | -
6 | Microscope. | 03/10/2017 | Not Needed | -
7 | Time with FBG Machine (Currently >=3 projects are using it). | 03/10/2017 | Resolved | -

## Slide 25
No. | Dependency | Resolve by | Status | Plan B
8 | Availability of tissues, membranes, and phantom for the eye surgery. | ~ 03/10/2017 | In Progress | -
9 | Getting time with surgeon to collect the data and validate the results the results obtained for the nonlinear function. | 05/10/2017 | Not started | Validate results with previously acquired (incomplete) data
Dependencies

## Slide 26
[1] Xingchi He, Force Sensing Augmented Robotic Assistance for Retinal Microsurgery, PhD Thesis, Jul 2015, Johns Hopkins University, Baltimore
[2] P. Gupta, P. Jensen, and E. de Juan, “Surgical forces and tactile perception during retinal microsurgery,” in International Conference on Medical Image Computing and Computer Assisted Intervention, vol. 1679, 1999, pp. 1218– 1225.
[3] S. Charles, “Techniques and tools for dissection of epiretinal membranes.”, Graefe’s archive for clinical and experimental ophthalmology = Albrecht von Graefes Archiv für klinische und experimentelle Ophthalmologie, vol. 241, no. 5, pp. 347–52, May 2003.
[4] R. Taylor, P. Jensen, L. Whitcomb, A. Barnes, R. Kumar, D. Stoianovici, P. Gupta, Z. Wang, E. DeJuan, and L. Kavoussi, “A Steady-Hand Robotic System for Microsurgical Augmentation,” The International Journal of Robotics Research, vol. 18, no. 12, pp. 1201–1210, 1999
[5] B. Mitchell, J. Koo, I. Iordachita, P. Kazanzides, A. Kapoor, J. Handa, G. Hager, and R. Taylor, “Development and application of a new steady-hand manipulator for retinal surgery,” in IEEE International Conference on Robotics and Automation, 2007, pp. 623–629.
Reading List

## Slide 27
[6] Candès, Emmanuel J., et al. "Robust principal component analysis?." Journal of the ACM (JACM) 58.3 (2011): 11.
[7] A. Menciassi, A. Eisinberg, G. Scalari, C. Anticoli, M. Carrozza, and P. Dario, “Force feedback-based microinstrument for measuring tissue properties and pulse in microsurgery,” in IEEE International Conference on Robotics and Automation, 2001, pp. 626–631.
[8] He X, Balicki M, Gehlbach P, Handa J, Taylor R, Iordachita I. A Multi-Function Force Sensing Instrument for Variable Admittance Robot Control in Retinal Microsurgery. IEEE International Conference on Robotics and Automation : ICRA : [proceedings] IEEE International Conference on Robotics and Automation. 2014;2014:1411-1418. doi:10.1109/ICRA.2014.6907037.
Reading List