## Slide 1: Force Control Algorithms for Sclera Eye Surgery
Team Members: Ankur Gupta, Saurabh Singh
Mentors: Dr. Iulian Iordachita, Dr. Marin Kobilarov and Dr. Russell H. Taylor
Date: 02/28/2017
Project - 20

## Slide 2: Motivation
The motivation for this project came from the challenges in the eye surgery :
Hand tremor of the surgeon
Excessive forces on the sclera leads to corneal striae
High risk of retinal tears (vision loss)
Inability of human to detect feeble forces
Eye surgery is difficult, requires intensive surgical training and practice
www.eyemdlink.com
Ftip
Fsclera
ds
Force Sensor
Force Sensor

## Slide 3: Background
Taken in part from
Xingchi He
Dissertation Defense
June 30th, 2015

## Slide 4: Background: Microsurgical Workstation
Eye Robot
Video microscope
Stereo display
Phantom eye and skull
Smart tools (Force/OCT)

## Slide 5: Background: Admittance Control
Constant Admittance Control
  Isotropic admittance
  The user input force applied on the instrument handle controls the velocity with which the robot follows the user motion.
The Solid lines in the above figure show the signal flow in current robot control algorithm, dashed lines show the signals that can also be incorporated into the control law. The constant admittance control is given by equations below.

## Slide 6: Background: Admittance Control
Constant Admittance Control
  Isotropic admittance
  Haptic feedback using force scaling
  The robot cooperative control law incorporates scaled sclera force measurement
Fss
Fhs

## Slide 7: Background: Variable Admittance Control
Isotropic admittance with force scaling
  Haptic feedback
  e.g. force scaling of sclera force
Anisotropic admittance to render virtual fixture
  Motion guidance/constraint
  e.g. virtual RCM
6-DOF motion
4-DOF motion
RCM

## Slide 8: Background: Variable Admittance Control
Isotropic admittance with force scaling
Anisotropic admittance to render virtual fixture
6-DOF motion
4-DOF motion
RCM

## Slide 9: Background: 2-DOF Force Sensing Tools
Fiber Bragg Grating (FBG sensors)
F
σ
Eye
FBG
FBGs
Fy
Fx
2-DOF Pick Tool
FBG sensors
He et al. SPIE Photonic West 2012
Expansion/contraction of FBG causes shift in wavelength. This driver and calibration has already been done and calibration matrices made available.
Notes: Previous work: a family of 2-DOF force sensing instrument (pick, forceps tools)
Explain FBG, How FBG works
Explain design concept of 2-DOF tools

## Slide 10: Background: Tool to Robot Calibration
Standard pivot calibration
  Tool shaft rigid
  Tool tip pivots around a fixed point
Pseudo Pivot Calibration
  Tool shaft flexible
  Tool shaft goes through a fixed point
  Use sclera force feedback to enforce RCM
Pivot Motion
RCM Point
{h}
{s}
{t}
Notes: ? Transformation from tool frame to robot frame
Pivoting the tool about a fixed insertion port with the robot
to prevent the tool from bending, we use force scaling with a nominal estimate of the transformation from the tool to the robot

## Slide 11: Overview
The transition function between the lower bound and the upper bound depth of the sclera depth is considered to be the straight line function but a study has to be conducted for the nonlinear functions as it is assumed that it won't be a linear function for human operative conditions.
Develop a learning based control algorithm. The learning is aimed to be, finally, performed by demonstration by the experienced surgeons.

## Slide 12: Detailed Plan

## Slide 13
Planned Timeline

## Slide 14: Deliverables
Minimum:
Procured software and Algorithm bridge source code
Ported backbone code to C++, if required
Procured and Calibrated Fiber-Optic Force sensing tool (Replace Fiber Optic tool, with low cost workable tool, if FO not procured)
Identified and documented source code that pertains to Sclera force sensing, drivers, APIs, and backbone algorithm
Restored the full setup to state, working with all tools and sensors attached.
Analysis of force feedback curve with multiple and unique coefficients

## Slide 15: Deliverables
Expected:
Introduction of non-linear behaviour to force feedback
Proof of concept with ‘untrained’ hands on Chorioallantoic membrane or synthetic sclera-impersonator membrane, such that forces seem natural in the process.

## Slide 16: Deliverables
Maximum:
Data collection from ‘experts’ or doctors
Validation of model with ground truth force feedbacks.

## Slide 17: Dependencies
No. | Dependency | Resolve by | Status | Plan B
1 | Develop complete understanding of problem and survey for possible approaches, based on feedback from mentors. | 3/3/2017 | In Progress | No Alternative
2 | Procure codebase | 02/24/2017 | Done | -
4 | Xingchi He’s time to discuss existing code and approach. | 3/3/2017 | In progress | Self Learning
5 | Fiber Optic sensor from Dr. Iordachita. | 3/3/2017 | Done | -
6 | Microscope. | 03/10/2017 | In progress | Borrow from other labs
7 | Time with FBG Machine (Currently >=3 projects are using it). | 03/10/2017 | Can only be scheduled after Ben Gorec’s Arrival | Work late night when no one is using the machine

## Slide 18: Dependencies
No. | Dependency | Resolve by | Status | Plan B
8 | Availability of tissues, membranes, and phantom for the eye surgery. | ~ 03/10/2017 | Will be done after Berk’s return | -
9 | Getting time with doctors to collect the data and validate the results the results obtained for the nonlinear function. | 05/10/2017 | Not started | Validate results with previously acquired (incomplete) data

## Slide 19: Division of Work
Saurabh | Ankur
(MechE and Controls background) | (Programming and ML background)
Restore the full setup to state, working with all tools and sensors attached. | 
Control Design | Crop the codebase to relevance
Calibration of tool | Port to C++
Acquisition of data from force sensor | 
Analysis of force feedback from multiple pose of tool wrt RCM | Experimenting with force feedback curve with multiple and unique coefficients
Coefficient Training and validation with non linear parameters | 
Integration of learned model with existing codebase | 

## Slide 20: Q&A