## Slide 1: Automatic Calibration of an FBG-based Force-Sensing Instrument for Intraocular Surgery
Nili Abtahi
Mentors: Prof. Iulian Iordachita, Mojtaba Esfandiari
April 2023

## Slide 2: Summery
Overview
Importance
Limitations

## Slide 3: Setup
Calibration and validation data collection:
Rotation angles: 0°, +90°
Tool shaft locations: tip, 12 - 20mm insertion depths at 2mm increments
±2g for tip force, ±10g everywhere else

## Slide 4
Project Plan & Methodology

## Slide 5: Setbacks
Uncategorized data
Challenge: initial data labeling and division
Current approach: gained experience for effective sample labeling
Suggested solution: update the code to categorize data based on angle, tip force and sclera force
Data set validation
Challenge: ensuring data set validity after testing
Current approach: reevaluate and restart if necessary
Suggested solution: use the mentioned categorization to validate for each part of data collection process

## Slide 6: Mathematical Framework
Courtesy of Dr. Changyan He.

## Slide 7: Project Flowchart

## Slide 8
Calibration Matrices:
Results & Analysis

## Slide 9: Result & Analysis
01
Tip Force Calibration
Calibrated
● FBG1 ● FBG4 ● FBG7
FBG-I
FBG-II
FBG-III
● FBG2 ● FBG5 ● FBG8
● FBG3 ● FBG6 ● FBG9
Matrix

## Slide 10: Results & Analysis
02
Tip Force Validation
Validated
Actual = calculated line

## Slide 11: Result & Analysis
03
Sclera Force Calibration
Calibrated
● FBG2 ● FBG5 ● FBG8
● FBG3 ● FBG6 ● FBG9
FBG-III
FBG-II
Matrix

## Slide 12: GUI & App
Reading information from FBG sensors
Selection of the applied load range
The command to start or stop the calibration operation
Important calibration parameters

## Slide 13: Revised Plan & Ideas
Linearity:
Until now, we have obtained an algebraic and first-order relationship between the FBG sensors and the force values ft, fs, which is a comprehensive and complete model because non-linear functions are not possible for FBGs.
Temperature :
Eliminate the effect of temperature changes on the FBG sensors by using methods such as neural network or other forms of analysis.
Factorial Experiments :
Factorial experiments can be used to reduce the amount of time required for calibration in this project by identifying the most important factors that affect the accuracy of the force-sensing instrument, and optimizing these factors using a minimum number of experiments.

## Slide 14: Deliverables
Minimum:
Experimental setup and documented calibration procedure
Automatic data analysis framework
Calibration of the force-sensing instrument
Validation of the results
Expected:
Software Requirements Specification (SRS)
GUI Design Documents
GUI Prototype
Source Code
Maximum:
Developing a mobile application for remote control
Minimum:
Experimental setup and documented calibration procedure
Automatic data analysis framework
Calibration of the force-sensing instrument
Validation of the results
Expected:
Software Requirements Specification (SRS)
GUI Design Documents
GUI Prototype
Source Code Mobile App
Maximum:
Nonlinear Calibration (higher order polynomials)
Temperature Compensation
Factorial Experiments

## Slide 15: Dependencies
No.​ | Dependency​ | Resolve by​ | Status​ | Plan B​
1​ | Eye-Robot 2.1​ | 02/01/2023​ | Available​ | No Alternative​
2​ | FBG Sensor Needle​ | 02/02/2023​ | Available​ | No Alternatives​
3​ | Calibration Setup​ [e.g. Scale, FBG Interrogator, Force Sensing Instrument]​ | 02/01/2023​ | Available ​ | Using the current tool​
4​ | Code Package​ | 02/01/2023​ | Available ​ | Failure to Accomplish Deliverables​
5​ | Mentors Feedback​ | Ongoing​ | Available | Delay in Delivering ​

## Slide 16: Timeline

## Slide 17: Procedure Documentation:
Ducument Title | Document Type | Status | Purpose | Author
Calibration Packag | Manual Instruction | Under Review | Step by step explanation of procedure to clarify the process | Nili Abtahi
Calibration Validation Procedure General | Flowchart | Final | Clarification and Implementation | Nili Abtahi
Tip Force Calibration Procedure | Flowchart | Final | Clarification and Implementation | Nili Abtahi
Sclera Force Procedure | Flowchart | Final | Clarification and Implementation | Nili Abtahi
Mathematical Framework | Scheme | Final | Clarification and Implementation | Changyan He
Matrix Calculation | Flowchart | Final | Clarification and Implementation | Nili Abtahi

## Slide 18: Management Plan
Project Meeting:
Weekly meetings with Prof. Iordachita & Mojtaba, Tuesdays at 5 p.m.
Lab Meeting:
Weekly meetings, Hackerman 306, Fridays at 4 p.m.
Mentor Collaboration & Guidance:
regular check-ins via email, phone calls & in-person with Mojtaba Esfandiari
Dr. He Consultation:
Through email for fact-checking, insights, and details verification

## Slide 19: Paper 2: Results & Conclusions
Sensitivity: 150 mN
Linearity: 99.4 %
Repeatability
Tip Force: 5 mN
Sclera Force: 10 mN
proposed calibration process provides a reliable and accurate method for ensuring the accuracy and reliability of force measurements!

## Slide 20: References
X. He, J. Handa, P. Gehlbach, R. Taylor, and I. Iordachita, “A Sub-Millimetric 3DOF Force Sensing Instrument with Integrated Fiber Bragg Grating for Retinal Microsurgery,” IEEE Transactions on Biomedical Engineering, vol. 61, no. 2, pp. 522 – 534, 2014.
X. He, M. Balicki, P. Gehlbach, J. Handa, R. Taylor, and I. Iordachita, “A MultiFunction Force Sensing Instrument with Variable Admittance Robot Control for Retinal Microsurgery,” IEEE International Conference on Robotics and Automation (ICRA), Hongkong, China, May 31-June 7, 2014.
I. I. Iordachita, M. D. De Smet, G. Naus, M. Mitsuishi and C. N. Riviere, "Robotic Assistance for Intraocular Microsurgery: Challenges and Perspectives," in Proceedings of the IEEE, vol. 110, no. 7, pp. 893-908, July 2022, doi: 10.1109/JPROC.2022.3169466.
Zhenglong Sun, M. Balicki, Jin Kang, J. Handa, R. Taylor and I. Iordachita, "Development and preliminary data of novel integrated optical micro-force sensing tools for retinal microsurgery," 2009 IEEE International Conference on Robotics and Automation, Kobe, Japan, 2009, pp. 1897-1902, doi: 10.1109/ROBOT.2009.5152836.

## Slide 21: Copyright The Johns Hopkins University 2022. All rights reserved.