## FMEA
Severity Scale:
1 | No harm to patient or operator and minimal disturbance to procedure
2 | Minimal harm/chance of harm to patient or operator (very minor injury or disturbance to procedure)
3 | Moderate harm/ chance of harm to patient or operator (minor injury or disturbance to procedure)
4 | Serious harm/chance of harm to patient or operator (serious injury or significant disturbance of procedure)
5 | Severe harm/chance of harm to patient or operator (life threatening or procedure failure)
Occurrence scale: |  |  |  |  | Detection scale:
1 | Extremely unlikely (should not occur during trials) |  |  |  | 1 | Detection certain (fault will always be detected)
2 | Minimal chance of occurrence (may occur once during whole set of trials) |  |  |  | 2 | Detection probable (fault is likely to be detected)
3 | Moderate chance of occurrence (may occur once in 10 procedures) |  |  |  | 3 | Detection possible (fault has approximatley 50% chance of being detected)
4 | Likely chance of occurrence (may occur once in every 2-3 procedures) |  |  |  | 4 | Detection unlikely (fault will probably be undetected)
5 | Certain occurrence for each procedure |  |  |  | 5 | Detection impossible (fault will never be detected)
Item/ Function | Potential Failure Mode | Potential Effects of Failure | S | Potential Cause(s) | How Failure is Detected | O | Current Controls | D | RPN (SxOxD) | Reccomended Actions
RoboELF | Unsanitary draping/undraping procedure | Spread of contamination between patients | 4 | Unsanitary draping/undraping procedure | Observation of torn drape or improper drape coverage, Observation of potentially contaminating splatter on robot | 2 | Draping and cleaning procedure; disposable parts | 1 | 8 | Follow proper draping and cleaning procedures
| Physical Injury to Patient or Operator | Robot falls on patient or operator | 4 | Incorrect Installation: Improper attachment to bedrail | Check stability in setup procedure | 1 | Check robot stability during setup; large margin for error in bedrail attachment system | 1 | 4 | Tighten joints as specified in manual; verify stability during setup
|  |  | 2 | Incorrect installation: Improper tightening of joint collars | Check tightening of joint collars in setup procedure | 2 | Check robot stability during setup; all joints that could move due to gravity have friction collars to prevent unintended motion | 1 | 4 | Tighten joints as specified in manual; verify stability during setup
|  |  | 3 | Incorrect maintenance: Improper tightening of friction collars | Check tightening of friction collars during setup and maintenance procedures | 1 | Proper maintenance and checks during setup | 1 | 3 | Tighten friction collars according to maintenance specification
|  | Robot becomes electrified | 3 | Broken wire due to fatigue or improper assembly | Check robot grounding during maintenance, Fuse on robot electronics will trip if fault occurs in use. Circuit breaker on power strip as well. | 1 | Proper maintenance and electrical fuse on robot, proper grounding during manufacturing | 1 | 3 | Stop. Move the robot out of the way. Remove the endoscope from the robot. Continue surgery manually without the robot.
| Robot Becomes Unresponsive(Active), Possible uncontrolled motion, Robot could drive itself or endoscope into patient or operator | Computer could not accurately determine robot position | 2 | Simultaneous failure of potentiometer and encoder, preventing cross-checking | Detected by motor controller motion error | 1 | Galil Overcurrent check | 2 | 4 | Stop. Move the robot out of the way. Remove the endoscope from the robot. Continue surgery manually without the robot.
|  |  | 2 | Encoder or potentiometer failure due to manufacture defect, fatigue or faulty system wiring | Detected by cross checking between encoders and potentiometers | 1 | Cross checking between encoders and potentiometers; the controller stops the robot and informs the user of the error. | 1 | 2 | Stop. Move the robot out of the way. Remove the endoscope from the robot. Continue surgery manually without the robot.
|  |  | 2 | Loss of Ethernet Connection because of physical break or software error | Detected by Watchdog Timer | 1 | Watchdog Timer | 1 | 2 | Attempt to reinitialize system. If reinitializes successfully, then resume. If not able to reinitialize, then stop. Move the robot out of the way. Remove the endoscope from the robot. Continue surgery manually without the robot.
|  | Motor could move without a user command | 2 | Motor malfunction due to manufacture defect, fatigue or faulty wiring | Detected by motor controller motion error or cross checking between encoders and potentiometers | 2 | Galil Overcurrent check, Cross checking between encoders and potentiometers | 1 | 4 | Stop. Move the robot out of the way. Remove the endoscope from the robot. Continue surgery manually without the robot.
|  |  | 2 | Joystick malfunction due to manufacture defect, fatigue or faulty wiring | Detected by checking for inconsistent joystick commands or user observation of uncommanded motion | 2 | Emergency stop button, checks for inconsistent joystick commands | 1 | 4 | Stop. Move the robot out of the way. Remove the endoscope from the robot. Continue surgery manually without the robot.
| Robot Becomes Unresponsive(Non-Active), No possible motion | PC unexpectedly becomes disabled | 1 | Internal Software error | Watchdog timer on Galil will stop system | 1 | Watchdog Timer | 1 | 1 | Stop. Move the robot out of the way. Remove the endoscope from the robot. Continue surgery manually without the robot.
|  |  | 1 | Power surge damages computer | Surge protector on power strip | 1 | Surge protector | 1 | 1 | Stop. Move the robot out of the way. Remove the endoscope from the robot. Continue surgery manually without the robot.
| Scope damage | Scope cord not slack enough and becomes caught on something | 3 | Incorrect installation | Check range of motion of robot before procedure | 1 | Detailed setup instructions | 1 | 3 | Correct training for setup and surgeon
| Robot damage | Robot runs into/tries to move past physical limits | 2 | Limit switch failure due to fatigue or manufacture defect | Detected by motor controller motion error | 2 | Galil Overcurrent check, Soft limits act as backup | 1 | 4 | Stop. Move the robot out of the way. Remove the endoscope from the robot. Continue surgery manually without the robot.
|  | Object blocking arm or internal jam in arm | 3 | Obstacles in range of motion, internal mechanical problem | Motor controller motion error | 1 | Follow setup instructions | 1 | 3 | Stop. Move the robot out of the way. Remove the endoscope from the robot. Continue surgery manually without the robot.

## TestPlan
Failure Mode | Test Method | Expected Results | Error Message | Actual Results | Date | Testers | Signatures
Draping/Cleaning
Water gets into motor controller enclosure and causes damage to controller | Pour water on all parts of robot to test seal | No water should get through seals | NA | No water penetrated seals | 2013-03-07 | Jonathan Kriss, Kevin Olds
Inadequate draping/undraping procedure | Spray drape with paint before removal | No paint should get through drape. | NA | No paint penetrated drape | 2013-01-30 | Jonathan Kriss, Kevin Olds
Drape tears | Attempt to tear drape during normal operations, adjust handles, scope, etc. | Drape should not tear during normal operation | NA | The drape did not tear under normal working conditions | 2013-01-30 | Jonathan Kriss, Kevin Olds
Joint/Arm Mechanical Failure
Passive arm joint clamp becomes loose | Loosen joint clamp | Friction Collar should stop motion. | NA | Friction Collars(Properly tightened) stop motion as expected | 2013-01-10 | Jonathan Kriss, Kevin Olds
Friction collars not tightened correctly | Loosen collars | Loose joint will be clearly noticable while trying to make adjustments during setup. | NA | Joints are able to fall is joint is loosened. Add check in set up instructions to verify tight enough friction collars | 2013-01-10 | Jonathan Kriss, Kevin Olds
Incorrect Attachment to bedrail | Loosen attachment to bedrail | Robot should be unstable but not fall | NA | Robot is slightly unstable but not in danger of falling in any way, as expected. | 2013-01-10 | Jonathan Kriss, Kevin Olds
Object blocking arm or internal jam in arm | Block joint with immovable object, do not damage robot. | Galil current/torque limits stop robot. | Error! Motor Error Detected! | Galil error stops motion and turns off motors, as expected | 2013-01-10 | Jonathan Kriss, Kevin Olds
Motor failure due to broken wire or short | Disconnect power from a motor during motion. | Robot will not move on that axis. | None | Motion Stops, no error, as expected | 2013-01-10 | Jonathan Kriss, Kevin Olds
Mechanical break between the motor and joint | Loosen pot attachment during operation then command motion. | PC pot/encoder check will fail and stop motion. Estop activated. | System Error! Encoder Failure | Axis A: Encoder Failure Error, as expected. Axis B: Encoder Failure Error, as expected. Axis C: Encoder Failure Error, as expected. | 2013-01-10 | Jonathan Kriss, Kevin Olds
Switch/Joystick Failure
Broken wire, short causes robot to become electrified | Test that case is properly grounded(Using multimeter) | All metal parts of the case should be electrically grounded. | NA | Case is properly earth-grounded. | 2013-01-10 | Jonathan Kriss, Kevin Olds
Limit switch mechanically fails during operation | Disconnect each limit switch in turn during operation | Software limits/physical limits stop motion. No apparent change in behavior | None | Software limits stop all axes, as expected. | 2013-01-10 | Jonathan Kriss, Kevin Olds
Limit switch mechanically fails before operation begins | Remove each limit switch in turn before operation | Physical limits should stop robot during calibration. Galil Motion Error should stop calibration | Calibration Fail | Motor Error during calibration, as expected | 2013-01-22 | Jonathan Kriss, Kevin Olds
Joystick mechanical failure: joystick cannot change switch settings | Remove switch assembly. | Robot should not respond to commands. Program should throw an error. | Joystick Error | Joystick Error, as expected. | 2013-01-10 | Jonathan Kriss, Kevin Olds
Joystick double switch failure | Simulate single switch failure(short switch) | Robot should not respond to commands. Program should throw an error. | Joystick Error | Joystick Error, as expected. | 2013-01-10 | Jonathan Kriss, Kevin Olds
Failure Mode | Test Method | Expected Results | Error Message | Actual Results | Date | Testers | Signatures
Potentiometer/Encoder Failure
Pot has electrical or mechaincal failure before operation begins | Disconnect wires from each pot in turn before operation | Calibration will fail or runtime checks will fail. Estop activated | Calibration Failed | Calibration Failed, as expected | 2013-01-22 | Jonathan Kriss, Kevin Olds
Pot has electrical or mechaincal failure during operation | Disconnect wires from each pot in turn during operation | Runtime checks will fail immediately. Estop activated | Encoder Error | Encoder Error, as expected | 2013-01-22 | Jonathan Kriss, Kevin Olds
Encoder has electrical or mechanical failure before operation begins | Disconnect each encoder in turn before operation | Calibration will fail or runtime checks will fail. Estop activated | Calibration Failed OR Encoder Error | Calibration Failed, as expected | 2013-01-22 | Jonathan Kriss, Kevin Olds
Encoder has electrical or mechanical failure during operation | Disconnect each encoder in turn during operation | Runtime checks will fail immediately. Estop activated | Encoder Error | Encoder Error, as expected | 2013-01-22 | Jonathan Kriss, Kevin Olds
Both encoder and pot have electrical or mechanical failure | Disconnect several combination of pots/encoders(same, different axes) during and before operation | Calibration will fail or runtime checks will fail. Estop activated | Calibration Failed OR Encoder Error | Calibration Failed, as expected | 2013-01-22 | Jonathan Kriss, Kevin Olds
Commuincation Failure
Estop gets disconnected from PC | Unplug USB cord before/during run. | Before: Program should give error on startup. During: Program should stop. Power will be disconnected from motors. | System Error! Estop Connection Failed | Estop Connection Error, as expected | 2013-01-10 | Jonathan Kriss, Kevin Olds
Ethernet failure: motor controller loses communication with PC | Disconnect ethernet cable during different operation states (Initialization, Calibration, Runtime) | Init: Should see error during startup. Calibration: Robot should complete last issued move command then stop. Runtime: Watchdog timer should stop robot from PC. Estop activated | Connection Error! Check Connections and Restart System | Robot stops and an error is thrown as expected | 2013-01-10 | Jonathan Kriss, Kevin Olds
Galil Controller crashes or loses power | Turn off Galil during operation. | Watchdog timer should stop robot from PC. Estop activated | Connection Error! Check Connections and Restart System | Connection Error, Estop, as expected. | 2013-01-10 | Jonathan Kriss, Kevin Olds
Computer crash due to power surge or other failure | Kill PC program process during different operation states. | Watchdog timer should stop robot from Galil. | NA | Galil enters Abort mode and stops all motion, as expected | 2013-01-10 | Jonathan Kriss, Kevin Olds