## Slide 1: Robo-ELF Software
Jonathan Kriss
JHU CiiS Lab

## Slide 2: System Overview
TaskManager
RobotTask
Robot Control
Safety Mechanisms
GalilController
cisst wrapper
GUI
PC
Galil
Watchdog Timer
Digital/Analog I/O
Robot
Motors/Encoders
Potentiometers
Limit Switches
Joysticks
EStop
Ethernet
USB

## Slide 3: Task-level Diagram
Task Manager
Connects mtsTask and Device objects
RobotTask
devGalilController Object
Robot Control functions
Safety checks
qtRobotDevice
QT widget
Displays robot state

## Slide 4: RobotTask Functional Description
GalilController object
ProvidesThroatRobot Interface
Safety checks
  Estop Connection
  PC Watchdog Timer
  Encoder/Pot checking
  Joystick valid input check
Control Functions

## Slide 5: qtRobotDevice Functional Description
RequiresThroatRobot Interface
Gets state information from robotTask
Displays position and switch state on GUI

## Slide 6: devGalilController Functional Description
Cisst GenericObject
cisst wrapper for Galil C++ API
Receives input from:
  Potentiometers(analog)
  Encoders(digital)
  Joystick controls(digital)
  Limit switches(digital)

## Slide 7: System Startup/Run/Shutdown
Create Tasks
Configure() / Startup() tasks
Run() Tasks
Kill() / Cleanup() Tasks
robotTask
qtRobotDevice
Init Galil
Start Watchdog
Build Enc/Pot tables
Safely shutdown robot and programs
Robot Control
Safety Checks
Error Handling/ Reporting
Close GUI
Connect to EStop

## Slide 8: RobotTask Run Loop
Robot Status Checks
Safety Checks OK
Log Error
No
Yes
Get Joystick Input
Valid Input
No
Critical Error
Yes
Estop
Print Error
System Restart Required to Continue
Robot Control
No
Yes
Invalid too long
(>.5s)
No
Yes
Stop Motion

## Slide 9: Safety Checks
Estop
  Cuts power to motors
  Activated by robotTask, Stop button
Watchdog Timer
  Ensures good connection between Galil and PC
  50ms interval, 75ms non-critical timeout, 125ms critical timeout
Encoder/Potentiometer checking
  Compares values to detect encoder failure
Software limits
  Prevents commanded movement outside of those limits
  Implemented with Galil’s built-in software limits
Joystick Input Check
  Two switches per axis on joystick input to check consistency

## Slide 10: Safety Errors
Critical Errors
Errors that could present a danger to the patient
Immediate stop of robot through Estop
Some errors can be fixed with a restart and recalibration
Non-Critical Errors
Errors that do not pose a danger to the patient or cause the system to stop working

## Slide 11: Safety Errors
Non-Critical Errors
Galil Command errors
  Limit switch hit, ect.
Non-critical watchdog timeout
Invalid Joystick input
  Under error threshold time
Critical Errors
Communication failure
  PC <-> Galil
  PC <- > Estop
Encoder/ Pot Error
Galil System Error
  Power error, Motion Error
Invalid Joystick input
  Over error threshold time

## Slide 12: Emergency Stop
Relay controllable by serial interface
Each run loop, check connection
EmergencyStop() function opens relays and cuts power to motors
Can also be activated manually with button

## Slide 13: Watchdog Timer
PC and Galil send handshake signal every Run() loop
  50ms period
Two timeout levels
  75ms non-critical
  125ms critical
Non-critical timeout logs error, no Estop
Critical timeout activates Estop
Galil program contained in Watchdog.dmc

## Slide 14: Maximum Effects of Timeout
Maximum uncontrolled travel distance:
  Axis A(Scope tip control)
    ~ 290°/s max motor speed = 36° / timeout
    ~ 90° total scope tip movement possible
  Axis B(Scope rotation)
    ~140° /s max motor speed = 18° / timeout
    18° total scope shaft rotation possible
  Axis C(Scope translation)
    ~4.25 in/s max motor speed = .5 in/ timeout
    .5 in total scope translation possible

## Slide 15: Encoder/Potentiometer Checking
Build tables of values for each on startup
  New values should always be monotonically increasing
During run, compare current values to values in tables using interpolation
If error > margin of error, stop robot
  Margin of error = step size of the table

## Slide 16: Joystick Switch Checking
Double switches on each joystick axis provide redundant sensing
The status of the joystick switches is constantly checked to ensure they are working properly
If an invalid switch state lasts for more than .5s, an error is thrown and the Estop activated

## Slide 17: Robot Control
Velocity function based on how long joystick is held down.
Linear acceleration from V_MIN to V_MAX
Different min/max values for each axis

## Slide 18: Velocity Function
Velocity
Time
V_MIN
V_MAX
AccelStartTime
AccelStopTime
1s

## Slide 19: Error Handling
Sources of errors
  GalilController
  Safety Checks in robotTask
All exceptions thrown up to high-level functions
All errors logged to cisstLog.txt
Safety/System failures generate simple error messages to user

## Slide 20: Galil Exception Classes
All exceptions thrown up to robotTask functions
  Only ExcpCommError and ExcpSystemError handled individually
RobotException
  Parent class for all exceptions
ExcpCommError
  Generated by SendCommand when Galil returns a timeout or command error
ExcpSystemError
  Generic System errors, non-critical
Classes not caught by name in robotTask
  ExcpMotionError, ExcpPowerError, ExcpMotorOff

## Slide 21: robotTask Exception Classes
EStopException
  Thrown by Estop functions if connection is lost
EncoderException
  Thrown by encoder/pot checking functions if error is detected. Also by startup calibration functions
WatchdogException
  Thrown by watchdog timer function if it times out
Joystick Exception
  Thrown by input function if readings are inconsistent

## Slide 22: Error Logging
robotTask set to ALLOW_ALL
Std::cout set to allow errors only
  Only displays error messages relevant to user
cisstLog.txt set to “VERBOSE” allowance
  Creates a more detailed log file

## Slide 23: Qt GUI
Displays information about current state
  Current position(red/green), encoder values(green), pot values(green)
  Limit switch state(blue)
Reads from state table in robotTask
No commands sent from GUI to robot